# Overview
This document’s purpose is to define the communication protocol used by the Intellus Controller firmware to send chamber data parameters to the Intellus Cloud service. This protocol structure is based upon the STOMP Protocol Specification (version 1.2). This is an open source protocol definition that was designed for asynchronous message passing between clients and servers.
The STOMP protocol is a frame based protocol that was modeled on HTTP. A frame consists of a command, a set of optional headers, and an optional body. STOMP is text based but also allows for the transmission of binary data messages.
The Intellus Cloud service will only be using a sub-set of all the different features defined in the STOMP protocol specification. For more information on this protocol, please review the specification documentation which is located at http://stomp.github.io/index.html.

# Configuration
The Intellus Controller will need to be configured to enable the pushing of chamber data to the cloud service. The following steps will define the configuration procedure.

## Create Web Account
The Intellus Cloud web application will provide a web page in which a new customer will be allowed to create a new web account. The web account will provide the user with the ability to enter their payment information and provide credentials for accessing the web application.
Upon completion of the web account creation, the customer will receive the Chamber configuration information via email. The customer will be required to enter this information into the Intellus Controller in order for the chamber to start sending data to the cloud application.

## Configure Intellus Controller
The firmware will need to provide the ability for the user to enter in the following values:
- Cloud web application IP address.
- Cloud web application Port number.
  - Default value for non-encrypted connection will be 61613.
  - Default value for encrypted connection will be 61614;
- Intellus Cloud customer code value.
  - Unique 8-character code used to identify the web account.

# Data Messages
The following sections define the different types of messages that are used to communicate with the Intellus Cloud service. The final byte of each message is a NULL (0) byte. The value of the “content-type” header, if present, does not include the NULL byte at the end of the message.

## Message Responses
The server application may respond to any request with the following response messages.
The ERROR message response is used to indicate the data request operation failed. The error message body will contain an error message which documents why the request failed.

Response Message
ERROR
receipt-id: <request receipt identifier>
content-type: text/plain
content-length: <error message length>
<error message>
NULL

If a server request includes a receipt identifier, the server will send back an acknowledgement response. This response will use the RECEIPT message and does not contain any data.

Response Message
RECEIPT
receipt-id: <request receipt identifier>
NULL

## Open Connection
The firmware will initially need to open a connection with the service. The initial command for opening the connection will be used to authenticate the chamber. The STOMP message will be required before any other message will be accepted by the server.

Request Message
STOMP
accept-version:1.2
login: <customer code>/<chamber serial number>
passcode: <Base 64 Hash>
NULL

Initially the service will be configured to allow plain text chamber login information. Eventually this information will be encoded with some algorithm and private key information. The firmware and the service will be the only applications that will know the encryption algorithm and secret key value.
The passcode algorithm consists of combining the customer code and chamber serial number values with the private key information. The hash value is then computed from this string. The passcode value is the hash value converted into a base-64 string.
If the server does not authenticate the request, it will respond with an error message. If the chamber is authenticated, the server will respond with a CONNECTED message. This response message will include a session identifier that will need to be included in all other messages to the service.

Response Message
CONNECTED
version: <version identifier>
session: <session identifier>
server:IntellusCloud/x.x.x
NULL

## Chamber Configuration
Once a connection has been established, the chamber will initially need to send the chamber configuration information. There are seven different configuration data strings that will be sent from the firmware. Each configuration string will consist of comma separated values followed by a newline character. The configuration strings will be sent in the following order.
### Tag Field Names
This configuration value will be sent as a comma separated list of the string field names. This information would normally be the first line of data in a standard CSV data file. It will be used by the server to determine what type of data is contained in each data column.
### Digital Output Types
This configuration value will be sent as a comma separated list of 47 integer values used to determine how each of the hardware output values are configured. This information will be used by the server to determine what outputs are configured as lights. Below is the enumerated list of possible configuration values for each output. Only the values marked as a light output will be treated as a light in the cloud application.

Value	Digital Output Types Description
0	Output Type Off [LIGHT]
1	Output Type Light On Off [LIGHT]
2	Output Type Light Dim [LIGHT]
3	Output Type Logic
4	Output Type Event
5	Output Type Dali Dim [LIGHT]
6	Output Type Timer

### Process Input Types
This configuration value will be sent as a comma separated list of 7 integer values used to determine how each of the hardware input values are configured. This information will be used by the server to determine what input values are available to be displayed by the cloud application.
The process input will be considered enabled anytime the value is not equal to the value representing the off state.

Value	Process Input Types Description
0	Sensor Type Off
1	Sensor Type TC (Not Used) [ENABLED]
2	Sensor Type MV [ENABLED]
3	Sensor Type Volts [ENABLED]
4	Sensor Type MA [ENABLED]
5	Sensor Type RTD 100 [ENABLED]

### RM Configuration Types
This configuration value will be sent as a comma separated list of 4 integer values used to determine if the set point values for the process inputs are enabled. This information will be used by the server to determine what set point values are available to be displayed by the cloud application.
The set point values for process inputs 1-3 will be considered enabled if the process input is enabled. The set point values for process inputs 4-7 will only be considered enabled if the configuration value is equal to the value representing the control loop state and the process input is enabled.

Value	RM Configuration Types Description
0	Slot Type Disabled
1	Slot Type Off
2	Slot Type Digital Input
3	Slot Type Digital Output
4	Slot Type Analog Output
5	Slot Type Control Loop [ENABLED]
6	Slot Type Monitor Input

### Input Name Values
This configuration value will be sent as a comma separated list of 7 string values used to determine the name value assigned to each process input channel. Each string value will consist of a 4-character string that may be padded with blank characters.
### Input Unit Values
This configuration value will be sent as a comma separated list of 7 string values used to determine the units of measure value assigned to each process input channel. Each string value will consist of a 3-character string that may be padded with blank characters.
### Alarm Values
This configuration value will be sent as a comma separated list of 11 integer values used to determine what alarms are currently enabled. A value of one will indicate the alarm is enabled and a value of zero will indicate the alarm is disabled. The table below identifies the order the alarm values are specified in the list.

Value	Alarm Description
0	eALARM_AIN_PV_1
1	eALARM_AIN_PV_2
2	eALARM_AIN_PV_3
3	eALARM_AIN_LIMIT_1
4	eALARM_AIN_PV_1_DEVIATION
5	eALARM_OVERRIDE_CLP_2_CH_A (Only used internally – Do not use)
6	eALARM_OVERRIDE_CLP_2_CH_B (Only used internally – Do not use)
7	eALARM_AIN_PV_4
8	eALARM_AIN_PV_5
9	eALARM_AIN_PV_6
10	eALARM_AIN_PV_7

The chamber will use the generic SEND command for making this request. The destination header value will indicate that this is a configuration message. The chamber should also specify a receipt identifier to receive an acknowledgment for this request. If the service does not receive this configuration request, then all data requests will fail.

Request Message
SEND
session:
destination:/config
receipt: <request receipt identifier>
content-type: text/plain
content-length: <data length>
<Tag Field Names Configuration>
<Digital Output Types Configuration>
<Process Input Types Configuration>
<RM Types Configuration>
<Input Names Configuration>
<Input Units Configuration>
<Alarms Configuration>
NULL

## Chamber Data
Once the chamber configuration data has been sent, the firmware will be ready to start sending data messages to the cloud application. The firmware will be configured to send one data message every minute to the cloud application.
The service will only send a response message if the firmware includes a receipt identifier in the request. If the data does not include a date/time value, then this information will be required in the message header. The destination header value will indicate that this is a data message and not a configuration message.

Request Message
SEND
session:
destination:/data
receipt: <request receipt identifier> optional
content-type: text/plain
content-length: <data length>
<CSV Data Row>
NULL

## Connection Shutdown
The chamber can gracefully shutdown the communication channel by sending the standard DISCONNECT message. The service will respond with a RECEIPT message which will indicate that the network socket can be closed.

Request Message
DISCONNECT
receipt: <request receipt identifier>
NULL