# Software Requirements Specification

# For

# Percival Intellus Gen II

NOTICE TO PERSONS RECEIVING THIS DOCUMENT

Watlow of Winona, MN, claims proprietary rights to the material herein disclosed.   It is supplied in confidence, without prejudice to any patent rights of Watlow Electric Manufacturing Company, and may not be reproduced or used to manufacture anything shown therein without written permission from Watlow.

| Revision | Date | Author | Description of Change |
| --- | --- | --- | --- |
| A | 9/23/15 | Cathy Wade | Initial document release |
| B | 1/14/16 | Cathy Wade | Updated timer description |
| C | 12/08/16 | Cathy Wade | See B-1 description |

Watlow Engineering		____________________________________________________________     (Required)

Signature			    Title		                 Date

Watlow Product Management	____________________________________________________________     (Required)

Signature			    Title		                 Date

Customer Engineering		____________________________________________________________     (Optional)

Signature			    Title		                 Date

REVISION LOG

| Revision | Date | Initials | Description of Changes |
| --- | --- | --- | --- |
| 1 | 8/27/14 | CLW | Initial Draft |
| 2 | 3/5/15 | CLW | Added Modbus and XML information |
| 3 | 4/6/15 | CLW | Added Ethernet, Bluetooth, Web Pages and an updated menu document. |
| 4 | 4/16/15 | CLW | Added the comments to the program operation to define that running programs cannot be edited and that when a program is stopped, the programmed step values are saved to the manual settings. |
| 5 | 5/5/15 | CLW | Modified cloud section to include new requirements for configuration information to be sent in the header. |
| 6 | 7/30/15 | CLW | Added new sources for the analog outputs. Did some customer requested changes to clarify parts. |
| A | 9/23/15 | CLW | Initial release to document control. |
| A-1 | 11/13/15 | CLW | Revised description of timer output operation per customer request |
| A-2 | 1/8/16 | CLW | Updated display loop to add reporting defrost active, requires changes to the embedded “Intellus Menus Specification.xlsx”. Added DI Door Duration and Bluetooth Enable tags to the XML tags, requires changes to the embedded “Complete XML Tag List.xlsx”. |
| B | 1/14/16 | CLW | Release to document control. This revision is going out with 18.47.86 but the changes in A-2 are actually post 18.47.86. |
| B-1 | 3/29/16 | CLW | Added DLI Input Source to the XML tags, requires changes to the embedded “Complete XML Tag List.xlsx”. Documented changes to the DLI programming web page, requires changes to the embedded “Web Page Screen Shots.docx”. Added selectable DLI Input Source feature description. Increased range of all propband parameters, changed defaults for the RH input related parameters, requires changes to the embedded “Intellus Menus Specification.xlsx” |
| C | 12/08/16 | CLW | Release to Document Control. This matches Released Revision 19.00.102 |

Numeric revisions, draft form - alpha revisions, controlled.

Contents

1.	Overlay Keypad	6

1.1	Key Definitions	6

1.1.1	Time Key	6

1.1.2	Inputs Key	6

1.1.3	Lights Key	6

1.1.4	Enter Key	6

1.1.5	Up/Down Arrow Keys	6

1.1.6	Program Key	7

1.1.7	Datalog Key	7

1.1.8	Diagnostic Key	7

1.1.9	Alarms Key	7

1.2	LED Definitions	7

1.2.1	Heat LED	7

1.2.2	Cool LED	7

1.2.3	Humidify LED	7

1.2.4	Dehumidify LED	7

1.2.5	Auxiliary LED	7

1.2.6	Program LED	8

1.2.7	Alarm LED	8

1.2.8	Communications LED	8

2.	Power Relay	8

3.	Alarm Buzzer	8

4.	High End Versus Low End	8

5.	Menu Definitions	8

6.	Analog Input	9

6.1	Input Names	9

6.2	Analog Input 1 (Temperature)	9

6.3	Analog Input 2 (Humidity)	9

6.4	Analog Input 3 (Auxiliary)	9

6.5	Analog Input 4 (RMC External Slot 1)	9

6.6	Analog Input 5 (RMC External Slot 2)	9

6.7	Analog Input 6 (RMC External Slot 3)	9

6.8	Analog Input 7 (RMC External Slot 4)	9

7.	Analog Input Offsets	10

7.1	Single Offsets	10

7.2	Multiple Offsets	10

8.	Control Loop Definitions	10

8.1	Control Methods	10

8.2	Setpoint Management	10

8.3	Control Loop 1 (Temperature, Heat/Cool)	10

8.3.1	Multiple PID Sets	11

8.4	Control Loop 2 (Humidity, Humidify/Dehumidify)	11

8.5	Control Loop 2 Humidify/Dehumidify Override	11

8.6	Control Loop 3 (Auxiliary, ChA/ChB)	11

8.7	Control Loop 4 (RMC External Slot 1, ChA/ChB)	11

8.8	Control Loop 5 (RMC External Slot 2, ChA/ChB)	11

8.9	Control Loop 6 (RMC External Slot 3, ChA/ChB)	12

8.10	Control Loop 7 (RMC External Slot 4, ChA/ChB)	12

9.	Digital Outputs	12

9.1	General Information	12

9.2	Light On/Off Type	12

9.3	Light Dimming Type	12

9.4	Logic Type	13

9.4.1	Logic Source PV 1-7, SP 1-7	13

9.4.2	Logic Source PV3 Light, SP 3 Light	13

9.4.3	Logic Source DIN 1-27	13

9.4.4	Logic Source Event 1-47	13

9.4.5	Logic Source Deviation Alarm or CLP 1-7 Soft Alarm	13

9.4.6	Logic Source All Lights Off	13

9.4.7	Logic Source Any Light On	13

9.5	Event Type	13

9.6	DALI Dimming Type	13

9.7	Timer Type	14

9.8	Any Light On Definition	14

9.9	Day Definition	14

9.10	Night Definition	14

9.11	Light Lifetime	14

10.	Digital Inputs	14

10.1	General Information	14

10.2	Defrost Terminate Type	15

10.3	Door Type	15

10.4	Trigger Type	15

11.	Analog Outputs	15

11.1	General Information	15

11.2	Retransmit PV Type	15

11.3	Retransmit SP Type	15

11.4	Retransmit Light Power	15

12.	Alarms	15

12.1	General Information	15

12.2	Analog Input Sensor Error	16

12.3	Analog Input Soft Alarm	16

12.4	Limit 1 Alarm	16

12.5	Limit 2 Alarm	16

12.6	Deviation Alarm	16

12.6.1	Deviation Alarm set to Non-Latching	16

12.6.2	Deviation Alarm set to Latching	16

12.6.3	Exceptions for deviation alarm occurrence	16

12.7	Light Lifetime Alarm	17

12.8	Door Open Alarm	18

13.	Defrost	18

13.1	Defrost Start	18

13.2	Defrost End	18

13.3	Defrost End Delay	18

13.4	Defrost Outputs	18

13.4.1	Defrost Disabled	18

13.4.2	Defrost Enabled, Off State	18

13.4.3	Defrost Enabled, Running State	18

13.4.4	Defrost Enabled, End Delay State	18

14.	Programs	18

14.1	Run Manual	18

14.2	Program Create/Edit Program	18

14.2.1	24 Hour Program Type (Ramping or Non-Ramping)	19

14.2.2	Elapsed Program Type (Ramping or Non-Ramping)	19

14.3	Program Run Program	19

14.3.1	24 Hour Non-Ramping Program Type	19

14.3.2	24 Hour Ramping Program Type	19

14.3.3	Elapsed Non-Ramping Program Type	19

14.3.4	Elapsed Ramping Program Type	20

14.4	Program Create/Edit Sequence	20

14.5	Program Run Sequence	20

14.5.1	24 Hour Non-Ramping Program Type	20

14.5.2	24 Hour Ramping Program Type	20

14.5.3	Elapsed Non-Ramping Program Type	20

14.5.4	24 Hour Non-Ramping Program Type	20

14.6	Program Edit Diurnal	21

14.7	Program Run Diurnal	21

14.8	Program Edit DLI	21

14.9	Program Run DLI	21

15.	WeatherEze	21

16.	XML Tags	22

16.1	Format of Data Returned	22

16.2	Encrypted XML Tags	22

16.3	HTTP and USB Command Strings	22

17.	Auto Configuration	22

18.	Datalogging	22

18.1	Low End RAM Datalog	23

18.2	High End SD Card Datalog	23

18.2.1	High End Datalog Parameter List	23

19.	Flashloading	24

19.1	Successful Flashload Process	24

19.2	Corrupt Application After Flashload	24

20.	Secondary Micro	24

20.1	Secondary Micro Communications	24

20.1.1	Intellus to Secondary Micro	24

20.1.2	Secondary Micro to Intellus	25

20.2	Secondary Micro Analog Input	25

20.3	Secondary Micro Buzzer Control	25

20.4	Secondary Micro Power Relay Control	25

20.5	Secondary Micro Reprogramming	25

21.	Security	25

21.1	Logging In and Out	25

21.2	Security Level 4	25

21.3	Security Level 3	25

21.4	Security Level 2	26

21.5	Security Level 1	26

21.6	Menu Passwords	26

21.7	Password Override	26

22.	DALI	26

22.1	Transmission Method Off	26

22.2	Transmission Method Broadcast	27

22.3	Transmission Method Direct	27

22.4	Power Sources	27

22.5	Programming a DALI ballast address	27

23.	Battery Backed RAM Parameters	27

24.	Modbus Building Automation System	27

25.	RM Communications	27

26.	USB Communications	28

27.	Ethernet	28

28.	Cloud Connection	28

29.	Bluetooth	29

30.	Web Pages	29

30.1	Web Page Security	29

30.1.1	Administrator Security	29

30.1.2	General User Security	30

31.	Email	30

31.1	Test	30

31.2	Cyclic	30

31.3	Change of State	31

32.	Real Time Clock	31

# Overlay Keypad

(Graphic provided for reference only.)

## Key Definitions

### Time Key

In the display loop pressing this key will move to the User\Time menu.

In the User\Time menu pressing this key will move back to the display loop.

### Inputs Key

In the display loop pressing this key will move to the User\Inputs menu.

In the display loop pressing this key and the Enter key for 3 seconds will move to the User\Offsets menu.

In the User\Inputs menu pressing this key will move back to the display loop.

### Lights Key

In the display loop pressing this key will move to the User\Outputs menu.

In the User\Outputs menu pressing this key will move back to the display loop.

### Enter Key

On power up pressing the Alarm key and the Enter key will enter Password Override Mode.

In the display loop pressing this key and the Inputs key for 3 seconds will move to the User\Offsets menu.

This key is used to accept parameter changes when editing a value.

### Up/Down Arrow Keys

These keys are used to change a parameter value when editing a value.

In the display loop pressing the Up Arrow key and the Enter key for 3 seconds will move to the password protected User\Global menu.

In the display loop pressing the Up Arrow key and the Diagnostics key for 3 seconds will move to the User\Service Diagnostics menu.

In the display loop pressing the Up Arrow key and the Down Arrow key for 3 seconds will move to the password protected Setup menus.

In the display loop pressing the Down Arrow key and the Alarms key for 3 seconds will move to the Login/Logout menu.

In the display loop pressing the Down Arrow key and the Diagnostics key for 3 seconds will move to the Firmware Download menu.

### Program Key

In the display loop pressing this key will move to the Programs menu.

In the Programs menu, if a program is being edited pressing this key will move to the edit selections.

Otherwise in the Programs menu pressing this key will move to the display loop.

When a string is being edited pressing this key will move through the string’s characters.

### Datalog Key

In the display loop pressing this key will move to the User\Datalog menu.

In the User\Datalog menu pressing this key will move to the display loop.

### Diagnostic Key

In the display loop pressing this key will move to the User\Diagnostics menu.

In the User\Diagnostics menu pressing this key will move to the display loop.

In the display loop pressing the Up Arrow key and the Diagnostics key for 3 seconds will move to the User\Service Diagnostics menu.

In the display loop pressing the Down Arrow key and the Diagnostics key for 3 seconds will move to the Firmware Download menu.

### Alarms Key

On power up pressing the Alarm key and the Enter key will enter Password Override Mode.

In the display loop pressing this key will move to the User\Alarms menu.

In the User\Alarms menu pressing this key will move to the display loop.

In the display loop pressing the Down Arrow key and the Alarms key for 3 seconds will move to the Login/Logout menu.

## LED Definitions

### Heat LED

The Heat LED will be on whenever the Heat B output is active, which is the Channel A output on Control Loop 1.

### Cool LED

The Cool LED will be on whenever the Cool output is active, which is the Channel B output on Control Loop 1.

### Humidify LED

The Humidify LED will be on whenever the Humidify output is active, which is the Channel A output on Control Loop 2.

### Dehumidify LED

The Dehumidify LED will be on whenever the Dehumidify output is active, which is the Channel B output on Control Loop 2.

### Auxiliary LED

The Auxiliary LED will be on whenever the Channel A output on Control Loop 3 is active.

### Program LED

The Program LED will be on whenever Program Mode is not Manual Mode.

The Program LED will flash whenever a program is being edited.

### Alarm LED

The Alarm LED will flash whenever any alarm is active.

### Communications LED

The Communications LED will be on whenever there is Ethernet communications activity.

# Power Relay

The power relay is used to connect line power to the chamber electrical system’s main power relay/contactor.

In the Intellus Gen II, both the Secondary Micro and the Intellus application need to be running correctly with no Limit alarms or errors to have the power relay engaged.

The following situations will cause the power relay to be disengaged:

- The communications fails between the Secondary Micro and the Intellus.

- The Secondary Micro detects a Limit 2 sensor error or Limit 2 High or Low Alarm state.

- The Intellus detects a Limit 1 sensor error or Limit 1 High or Low Alarm state.

# Alarm Buzzer

The alarm buzzer is used to indicate an alarm condition.

Any of the defined alarms or errors will cause the buzzer to be activated when the alarm or error occurs.

Any key press when the buzzer is active will silence the buzzer.

If an alarm is active, the buzzer has been silenced and another alarm occurs, the buzzer will become active again.

The Secondary Micro and the Intellus application can separately control the buzzer output.

# High End Versus Low End

There are two possible configurations of the Intellus Gen II controller, the High End control has a part number ending in -0004 and the Low End control has a part number ending in -0003.

The High End control has the hardware installed to enable the following features:

- Ethernet

- Datalogging

- Web Pages

- SD Card

- Bluetooth

- Cloud

- DALI

The Low End control has a subset of Datalogging that stores information in a RAM buffer for the tablet to use.

# Menu Definitions

The menus are defined in the attached spreadsheet.

# Analog Input

## Input Names

All analog inputs have a programmable input name up to four characters.

## Analog Input 1 (Temperature)

This input is always used for temperature sensing using a 100 ohm RTD sensor and it is always enabled.

The units for this input follow the display units setting, either F or C.

The default input name is TEMP.

The default input units are C.

This input is the source for the Limit 1 Alarm.

## Analog Input 2 (Humidity)

This input is commonly used for humidity sensing but can be used for any process input of the allowed types.

The sensor types allowed for this input are Off, Volts and mA, default is Off.

The default input name is RH.

The default units name is RH%.

## Analog Input 3 (Auxiliary)

This input is commonly used for auxiliary sensing and can be RTD or any process input of the allowed types.

The sensor types allowed for this input are Off, Volts, mV and 100 ohm RTD, default is Off.

The default input name is AUX.

The default units name is PPM.

## Analog Input 4 (RMC External Slot 1)

This input is only available if an external RMC with an input in Slot 1 is attached.

The sensor types allowed for this input are Off, Volts, mV, mA and 100 ohm RTD, default is Off.

The default input name is RMC4.

The default units name is UML.

## Analog Input 5 (RMC External Slot 2)

This input is only available if an external RMC with an input in Slot 2 is attached.

The sensor types allowed for this input are Off, Volts, mV, mA and 100 ohm RTD, default is Off.

The default input name is RMC5.

The default units name is UML.

## Analog Input 6 (RMC External Slot 3)

This input is only available if an external RMC with an input in Slot 3 is attached.

The sensor types allowed for this input are Off, Volts, mV, mA and 100 ohm RTD, default is Off.

The default input name is RMC6.

The default units name is UML.

## Analog Input 7 (RMC External Slot 4)

This input is only available if an external RMC with an input in Slot 4 is attached.

The sensor types allowed for this input are Off, Volts, mV, mA and 100 ohm RTD, default is Off.

The default input name is RMC7.

The default units name is UML.

## Independent Temperature Limit Input

This input is always used for temperature sensing using a 100 ohm RTD sensor and it is always enabled.

The units for this input inherit the controller’s global setting for temperature units of either F or C.

The default input units are C.

This input is handled by the secondary micro and is the source for the Limit 2 Alarm

# Analog Input Offsets

All Analog Inputs have user adjustable offsets which are used to modify the calculated process value for calibration purposes.  The process value displayed will be the calculated process value plus the offset.

Analog Input 1 has the option to use the single offsets or to use multiple offsets.

## Single Offsets

All Analog Inputs have two separate offsets, one is defined to be used during the Day and the other during the Night.

The Day offset is used when the Day Definition (as stated under the Digital Outputs section) is true.  Otherwise the Night offset is used.

## Multiple Offsets

Analog Input 1 can be configured to use multiple offsets.

There are three Offset Values and three Offset Points each for Day and Night.

The Offset Points are process values.

When Analog Input 1 has a process value less than or equal to the Offset Point 1, the Offset Value 1 is used.

When Analog Input 1 has a process value greater than or equal to the Offset Point 3, the Offset Value 3 is used.

When Analog Input 1 has a process value equal to the Offset Point 2, the Offset Value 2 is used.

When Analog Input 1 has a process value greater than the Offset Point 1 and less than the first Offset Point 2 or greater than the Offset Point 2 and less than the Offset Point 3, the actual offset is calculated linearly over the Offset Point range using the Offset Value range.

# Control Loop Definitions

## Control Methods

The functionality of the control loop control uses the standard EZ-Zone libraries.

Each control loop can have separate control methods for Channel A and Channel B and can select between Off, PID and On/Off.

In addition, Control Loop 1 Channel B (Cool) can select Complementary.

## Setpoint Management

In Program Manual mode with WeatherEze disabled, the manual closed loop setpoint is used for all the control loops.

In Program Manual mode with WeatherEze enabled, the closed loop setpoint is expected to be continuously by the WeatherEze program.

In Program Run Program and Program Run Sequence modes, the programmed setpoint for the currently running program step is used for all the control loops.

In Program Run Diurnal, if the time is determined to be during the day, the programmed Diurnal Day setpoint is used, otherwise the programmed Diurnal Night setpoint is used.

In Program Run DLI, if the time is determined to be during the day, the programmed DLI Day setpoint is used, otherwise the programmed DLI Night setpoint is used.

## Control Loop 1 (Temperature, Heat/Cool)

The Control Loop 1 uses Analog Input 1 which is always enabled and set to 100 ohm RTD.

The Channel A and Channel B outputs are always enabled although the control methods can individually be set to Off.

Control Loop 1 has a special control method which can be used for Channel B (Cool) called Complementary.  If this method is chosen, the state of the Channel B output is controlled at the hardware driver level to be on when the Channel A output is Off and to be Off when the Channel A output is On.

Control Loop 1 also has a secondary heat output (Heat Output A) which is also controlled at the hardware driver level to follow the Channel A output, but to be overridden to Off if Control Loop 2 (Humidity) Channel B (Dehumidify) is On, and to be overridden to On if the Defrost B output is On.

If the control method is set to PID, the control outputs normally operate as Fixed Time Base using the Cycle Time.  If the Cycle Time is set to 0, the outputs operate as Variable Time Base.

### Multiple PID Sets

Multiple PID sets are only implemented for Control Loop 1 and can be overridden by setting all four sets of parameters to the same values.

The four PID sets are used depending on the Control Loop 2 (Humidity) Channel B (Dehumidify) output and Any Light On or All Lights Off.

PID set 1 is used for Any Light On and Dehumidify Off.

PID Set 2 is used for Any Light On and Dehumidify On.

PID Set 3 is used for All Lights Off and Dehumidify Off.

PID Set 4 is used for All Lights Off and Dehumidify On.

## Control Loop 2 (Humidity, Humidify/Dehumidify)

The control loop 2 uses Analog Input 2.

The Channel A and Channel B outputs default to disabled.

The control of the outputs follows a standard control loop if the outputs are enabled and the control method is not set to Off.

If the control method is set to PID, the control outputs normally operate as Fixed Time Base using the Cycle Time.  If the Cycle Time is set to 0, the outputs operate as Variable Time Base.

## Control Loop 2 Humidify/Dehumidify Override

The Control Loop 2 outputs can be overridden to Off by individually choosing an override source for each output.

The selections for the override source are Off, Input 1 Process Value (PV1), Input 1 Set Point (SP1), Digital Input 1 (DIN1), Digital Input 2 (DIN2), and Digital Input 3 (DIN3).

For the PV1 and SP1 selections, the action occurs much like a normal alarm.  If the override side is set to Both or High and the source is above the high setpoint, the output is turned off.  Likewise, if the override side is set to Both or Low and the source is below the low setpoint, the output is turned off.

For the DIN selections, the output is turned off if the input is active.

## Control Loop 3 (Auxiliary, ChA/ChB)

The control loop 3 uses Analog Input 3.

The Channel A and Channel B outputs default to disabled.

The control of the outputs follows a standard control loop if the outputs are enabled and the control method is not set to Off.

If the control method is set to PID, the control outputs normally operate as Fixed Time Base using the Cycle Time.  If the Cycle Time is set to 0, the outputs operate as Variable Time Base.

## Control Loop 4 (RMC External Slot 1, ChA/ChB)

This control loop is only available if an external RMC with a control loop in Slot 1 is attached.

The control loop 4 uses Analog Input 4.

The Channel A and Channel B outputs default to disabled.

The control of the outputs follows a standard control loop if the outputs are enabled and the control method is not set to Off.

If the control method is set to PID, the cycle method can be between Fixed Time Base using the Cycle Time and Variable Time Base.

## Control Loop 5 (RMC External Slot 2, ChA/ChB)

This control loop is only available if an external RMC with a control loop in Slot 2 is attached.

The control loop 5 uses Analog Input 5.

The Channel A and Channel B outputs default to disabled.

The control of the outputs follows a standard control loop if the outputs are enabled and the control method is not set to Off.

If the control method is set to PID, the cycle method can be between Fixed Time Base using the Cycle Time and Variable Time Base.

## Control Loop 6 (RMC External Slot 3, ChA/ChB)

This control loop is only available if an external RMC with a control loop in Slot 3 is attached.

The control loop 6 uses Analog Input 6.

The Channel A and Channel B outputs default to disabled.

The control of the outputs follows a standard control loop if the outputs are enabled and the control method is not set to Off.

If the control method is set to PID, the cycle method can be between Fixed Time Base using the Cycle Time and Variable Time Base.

## Control Loop 7 (RMC External Slot 4, ChA/ChB)

This control loop is only available if an external RMC with a control loop in Slot 4 is attached.

The control loop 7 uses Analog Input 7.

The Channel A and Channel B outputs default to disabled.

The control of the outputs follows a standard control loop if the outputs are enabled and the control method is not set to Off.

If the control method is set to PID, the cycle method can be between Fixed Time Base using the Cycle Time and Variable Time Base.

# Digital Outputs

## General Information

There are 23 digital outputs on the Intellus board and a possible additional 24 outputs on an external RME.

Outputs 24-29 are only available if an external RME with digital I/O in Slot 1 is attached and the slot is configured to be digital outputs.

Outputs 30-35 are only available if an external RME with digital I/O in Slot 2 is attached and the slot is configured to be digital outputs.

Outputs 36-41 are only available if an external RME with digital I/O in Slot 3 is attached and the slot is configured to be digital outputs.

Outputs 42-47 are only available if an external RME with digital I/O in Slot 4 is attached and the slot is configured to be digital outputs.

The output type can be selected from Off, Light On/Off, Light Dimming, Logic, Event, DALI Dimming or Timer.

## Light On/Off Type

This type controls the output either full on or full off.

During Program Manual Mode the light setting is controlled by the manual light on/off.

During Program Run Program or Program Run Sequence the light setting is controlled by the programmed light on/off in the currently running step.

During Program Run Diurnal and Program Run DLI the light setting is full on during the day and full off during the night.

If Defrost is running the light setting is full off.

In the display loop, an output of this type is represented by a 0 or a 1 in any prompt that has a output list and ON or OFF in the individual output prompt.

## Light Dimming Type

This type controls the output using a percentage which controls the output using Variable Time Base.

During Program Manual Mode the light setting is controlled by the manual light dimming percentage.

During Program Run Program or Program Run Sequence the light setting is controlled by the programmed light dimming percentage in the currently running step.

During Program Run Diurnal and Program Run DLI the light setting is full on during the day (100%) and full off during the night (0%).

If Defrost is running the light setting is full off (0%).

In the display loop, an output of this type is represented by a bar in any prompt that has a light list and XXX% in the individual light prompt.  The bar’s height is one line of dots for each 16.66% power level.

## Logic Type

The logic type output has a source selection of Off, PV1-7, SP 1-7, PV3 Light, SP 3 Light, DIN 1-27, Event 1-47, Deviation Alarm, All Lights Off, Any Light On and CLP 1-7 Soft Alarm.

The logic type output also has an output active level, High or Low.

In the display loop, an output of this type is not displayed.

### Logic Source PV 1-7, SP 1-7

For the PV and SP selections, the action occurs much like a normal alarm.  If the logic side is set to Both or High and the source is above the high setpoint, the logic output is turned on.  Likewise, if the logic side is set to Both or Low and the source is below the low setpoint, the output is turned on.

### Logic Source PV3-7 Light, SP 3-7 Light

These selections operate the same as the standard source for PV or SP except that the outputs are considered lights.  This means that they can be overridden to Off if Defrost is running or if the Program Run Diurnal or Program Run DLI is running and it is night.

### Logic Source DIN 1-27

For the DIN sources the logic output is active if the DIN source is active and the Digital Input Type is set to Trigger.

### Logic Source Event 1-47

The word Event in this type of source is not quite accurate anymore since the Gen II Intellus allows any digital output to be any type.  The more accurate label would be Digital Output for Event which will be used here.

For the Digital Output source the operation depends on what type of output the selected source has been defined as.  The only allowed output types that can be a logic source are Light On/Off, Event, Light Dimming or DALI Dimming.

For a Digital Output source of type Light On/Off or Event, the logic output is active when the source output is active.

For a Digital Output source or type Light Dimming or DALI Diming, there is an additional parameter called the Dimming Threshold.  The logic output is active when the source output percentage is above the Dimming Threshold.

### Logic Source Deviation Alarm or CLP 1-7 Soft Alarm

For the alarm sources the logic output is active if the selected alarm is active.

### Logic Source All Lights Off

For the All Lights Off source the logic output is active if the Any Lights On criteria is FALSE.

### Logic Source Any Light On

For the Any Light On source the logic output is active if the Any Lights On criteria is TRUE.

## Event Type

This type controls the output either full on or full off.

In the display loop, an output of this type is represented by a 0 or a 1 in any prompt that has an output list and ON or OFF in the individual output prompt.

## DALI Dimming Type

The DALI Dimming output type is exactly the same as Light Dimming Type except that the light output percentage is sent over the DALI interface instead of being produced on the physical digital output.

## Timer Type

This type controls the output either full on or full off based the cycle on and cycle off durations and the Active Temperature parameter.

Example:  Cycle on duration = 60, Cycle off duration = 30, Active Side = Low, Active Temp = 20C.  When the temp process is below 20C the output cycles on for 60 minutes, then off for 30 minutes, then on for 60 minutes and continually cycles.  When the temp is >= 20C the current timer accumulating on timer or off time is deactivated and reset until the temp goes back under 20C.  Timer deactivation means that the output is disabled, regardless of the state it was in before the timer output was disabled.  Once the temperature goes back under 20C, the timer output is energized and the timer starts accumulating time against the cycle on duration.  In the display loop, an output of this type is not displayed.

## Any Light On Definition

Any one of the following would cause Any Light On to be TRUE.

- Digital Output type set to Light On/Off and output is active.

- Digital Output type set to Light Dimming or DALI Dimming and the output is not a source to a logic output and the output percentage is greater than 0.

- Digital Output type set to Logic with an Event source and source output type is set to Light Dimming or DALI Dimming and source output percentage is greater than the dimming threshold.

- Digital Output type is set to Logic with PV3 or SP 3 Light source and the output is active.

## Day Definition

Day is defined to be when any output that would cause Any Light On to be TRUE is active.

The Gen I specification stated separately that during Program Mode Run Diurnal, Day is the time between the Day Period Start and Day Period End but since all light outputs are defined to be on during this time for Run Diurnal, the previous statement is sufficient.

The Gen I specification stated separately that during Program Mode Run DLI, Day is the time between the Day Period Start and Day Period End AND the accrued light integration is less than the Daily Light Integral but since all light outputs are defined to be on during this time for Run DLI, the previous statement is sufficient.

## Night Definition

Night is defined as any time when it is not Day.

## Light Lifetime

Any output that would cause Any Light On to be TRUE will increment the light lifetime.  This value is stored in nonvolatile only after at least an hour has accrued to avoid wearing out the memory.

# Digital Inputs

## General Information

There are 3 digital inputs on the Intellus board and a possible additional 24 inputs on an external RME.

Inputs 4-9 are only available if an external RME with digital I/O in Slot 1 is attached and the slot is configured to be digital inputs.

Inputs 10-15 are only available if an external RME with digital I/O in Slot 2 is attached and the slot is configured to be digital inputs.

Inputs 16-21 are only available if an external RME with digital I/O in Slot 3 is attached and the slot is configured to be digital inputs.

Inputs 22-27 are only available if an external RME with digital I/O in Slot 4 is attached and the slot is configured to be digital inputs.

The input type can be selected from Off, Defrost Terminate, Door or Trigger.  Multiple inputs can be set to the same type.

## Defrost Terminate Type

This type terminates an active defrost if the input becomes active.

Any input set to this type can terminate an active defrost cycle.

## Door Type

This type produces an alarm and door open display if the input becomes active and stays active for the programmed door duration.  Setting the door duration to 0 will trigger an immediate display and alarm.  Pressing any key will acknowledge the door alarm.

If multiple inputs are set to be door types, a new input becoming active after the door alarm has been acknowledged will retrigger the door alarm.

## Trigger Type

This type can be used as a Digital Output Logic Source or as Control Loop 2 Humidify/Dehumidify Override Source (input 1-3 only).

If it is used as a Digital Output Logic Source, the logic output is active if the input is active.

If it is used as an Override Source, the Control Output is inactive if the input is active.

# Analog Outputs

## General Information

There are no analog outputs on the Intellus board but a possible 12 outputs on an external RME.

Outputs 1-3 are only available if an external RME with analog outputs in Slot 1 is attached and the slot is configured to be analog outputs.

Outputs 4-6 are only available if an external RME with analog outputs in Slot 2 is attached and the slot is configured to be analog outputs.

Outputs 7-9 are only available if an external RME with analog outputs in Slot 3 is attached and the slot is configured to be analog outputs.

Outputs 10-12 are only available if an external RME with analog outputs in Slot 4 is attached and the slot is configured to be analog outputs.

The output type can be selected from Off, Retransmit PV, Retransmit SP, Retransmit Light Power, Retransmit Heat Power or Retransmit Cool Power.

The analog output electrical type can be selected from Off, mA or Volts.

## Retransmit PV Type

This type is used to retransmit the current process value from the selected Analog Input.

## Retransmit SP Type

This type is used to retransmit the current setpoint from the selected Control Loop.

## Retransmit Light Power

This type is used to retransmit the current light power level from the selected Digital Output.

## Retransmit Heat Power

This type is used to retransmit the current heat power level (Channel A) from the selected Control Loop.

## Retransmit Cool Power

This type is used to retransmit the current cool power level (Channel B) from the selected Control Loop.

# Alarms

## General Information

There are 12 possible alarm sources on the Intellus board and 6 possible sources from an external RMC.

Each of the three local Analog Inputs or the four external Analog Inputs can have a sensor error or a soft alarm.

There is an independent input on the secondary micro which is used as a Limit.  This input can have a sensor error or a Limit 2 alarm.

The first Analog Input has a Deviation alarm and a Limit 1 alarm.

There are also the Light Lifetime alarm and the Door Open alarm.

An active alarm causes the buzzer to occur and moves to the appropriate alarm display in the display loop.  The buzzer can be silenced by a button press.

If an alarm has been silenced and there is another alarm, the buzzer will occur again.  The new alarm will be part of the display loop but it may not appear without scrolling through the loop.

The power relay will also be deactivated if either the Limit 1 or Limit 2 has a sensor error, High or Low Alarm.

## Analog Input Sensor Error

For any analog input whose sensor type is set to RTD 100 ohm, the sensor error is produced if the input is shorted or open.  This alarm occurs for any analog input set to RTD 100 ohm and cannot be disabled.

For any analog input whose sensor type is one of the process types (mA, Volts, mV), the sensor error is produced using the Process Error settings.  This alarm can be disabled by setting the Process Error Enable to No.  If the alarm is enabled and the Analog Input Process Value goes below the Process Value Low, the alarm will occur.

## Analog Input Soft Alarm

For an analog input whose sensor type is set to a type other than Off this alarm can be enabled.  This alarm becomes active if it is enabled and the Analog Input Process Value goes outside the range defined by the High and Low values and the Side selection (Both, High, Low).

## Limit 1 Alarm

This alarm cannot be disabled.  It uses the Analog Input Process Value from Analog Input 1 and becomes active if the Analog Input Process Value goes outside the range defined by the High and Low values (Side is always Both).

If this alarm becomes active or there is a sensor error on Analog Input 1 the power relay is deactivated.

## Limit 2 Alarm

This alarm cannot be disabled.  It uses the redundant input on the secondary micro and becomes active if the input PV goes outside the range defined by the High and Low values (Side is always Both).

If this alarm becomes active or there is a sensor error on the redundant input the power relay is deactivated.

## Deviation Alarm

This alarm uses the Analog Input Process Value from Analog Input 1 and becomes active, with exceptions as noted below, when the input PV goes outside the range defined as the number of degrees above and below the Control Loop 1 Setpoint or the range of normal operation (Side is always Both).

The deviation alarm can be disabled.

The deviation alarm audible can be disabled.

The deviation alarm can be set to non-latching or latching.

### Deviation Alarm set to Non-Latching

If the alarm is set to non-latching, the alarm is cleared when the input PV returns to the range of normal operation.

### Deviation Alarm set to Latching

If the alarm is set to latching, any key press after the alarm has been silenced and the input PV reenters the range of normal operation will clear the alarm.

If the alarm is set to latching, if a deviation alarm has been silenced but not cleared, when the input PV reenters the range of normal operation and then exits it again, the audible will reactivate.

### Exceptions for deviation alarm occurrence

If a deviation condition occurs because of power up or setpoint change, the alarm will not occur until the time defined by the Deviation Silence has elapsed or the input PV has entered the range of normal operation and then goes outside the range.  Note that setting the Deviation Silence to 0 minutes causes an immediate alarm in this case.

If the Intellus is power cycled while a deviation alarm is active or latched, the alarm will reoccur on power up whether or not the input PV is within the range of normal operation.

The following diagrams show deviation alarm action during a running program, note that Ramping Mode means either 24 Hour Ramping or Elapsed Ramping program types and Non-Ramping Mode means either 24 Hour Non-Ramping or Elapsed Non-Ramping program types.

Clarification on the deviation alarm operation during programs.  For 24 Hour Ramping, the deviation silence time is only used for the intial step transition.  For Elapsed Ramping it is not used at all.  For all Non-Ramping the deviation silence time is used for each setpoint change.

## Light Lifetime Alarm

This alarm will occur when any output’s Light Lifetime accrued time is greater than or equal to the Light Life Alarm time.

If the alarm has been silenced and another output’s Light Lifetime accrued time is greater than or equal to the Light Life Alarm time, the buzzer will restart.

Any output that would cause Any Light On to be TRUE will increment the light lifetime, note that the time is saved in non-volatile memory every hour.

The non-volatile memory value allows the alarm to reoccur through a power cycle.

The Light Lifetime can be reset to 0 by the user, which is the only way to totally clear the alarm.

## Door Open Alarm

This alarm will occur when any Digital Input is set to a Door Type and the input becomes active.

If multiple inputs are set to be door types, a new input becoming active after the door alarm has been acknowledged will retrigger the door alarm.

# Defrost

The Defrost function can be disabled.

If the Defrost function is enabled the number of defrost cycles, start time, duration and end delay can be set.

## Defrost Start

Defrost will become active if it is enabled, the time of day reaches the programmed start time and the Analog Input 1 Process Value is less than the programmed Defrost Active Temperature.

## Defrost End

Defrost will end if the time of day reaches the time calculated by adding the programmed start time to the duration.

Defrost will end if any Digital Input is set to type Defrost Terminate and the input is active.

## Defrost End Delay

The Defrost End Delay is used to delay the turning on of the Defrost output A from the time that the Defrost has ended.

## Defrost Outputs

### Defrost Disabled

Defrost output A is turned off and output B is turned off.

### Defrost Enabled, Off State

Defrost output A is turned on and output B is turned off.

### Defrost Enabled, Running State

Defrost output A is turned off and output B is turned on.

All Digital Outputs that would cause Any Light On to be TRUE will be forced off.

Control output Heat A is turned on.

All control outputs in Control Loop 1 and Control Loop 2 are turned off.

Control Loops 2-7 operate normally.

### Defrost Enabled, End Delay State

Defrost output A is turned off and output B is turned off.

# Programs

## Run Manual

The Manual program mode uses the manual setpoints for the control loops and the manual output settings for the digital outputs if the WeatherEze is disabled, otherwise it gets the control information from the WeatherEze program.

When the program mode is changed from any running program to Run Manual, the current programmed step values will be saved to the manual setpoints for the control loops and the manual output settings for the digital outputs.

## Program Create/Edit Program

Edit Program will not be available if no program has been created.

A program cannot be edited if it is running.

All programs have a name and a number of steps which contain time information, the Control Loop setpoints and the lighting settings.

No two programs can have the same name.

Only digital outputs that are set to Light On/Off, Event On/Off, Light Dimming or DALI Dimming can be edited in a program step.

### 24 Hour Program Type (Ramping or Non-Ramping)

There can be up to 50 24 Hour Programs created with a total of 500 steps divided among them.

The start time parameter in a 24 Hour program step is a time of day.

No two steps can have the same time of day.

After a new step has been added, the program will be reordered so that the steps are in chronological order.

### Elapsed Program Type (Ramping or Non-Ramping)

There can be up to 20 Elapsed Programs created with a total of 60 steps divided among them.

The time parameter in an Elapsed program step is the duration of the step.

New steps are added after the last step.

## Program Run Program

This mode will not be available if there is no program created.

During Program Run Program any control loop setpoint is controlled by the programmed level in the currently running step.

During Program Run Program any light setting is controlled by the programmed level in the currently running step.

A running program cannot be edited.

Once a program is started, it runs continuously until the Program Mode is changed.

The Run Program, current program and current step are saved through a power cycle.

When a single program finishes its last step, it begins again at the first step and continues this cycle until the Program Mode is changed.

### 24 Hour Non-Ramping Program Type

The start time parameter in a 24 Hour program step is a time of day to begin using the programmed setpoints and light settings.

The end time is the start time parameter in the next chronological step.

The 24 Hour Non-Ramping type uses the programmed setpoints and light settings for the duration of the currently running step.

When a 24 Hour program is started, the first step is the one where the current time falls within the step duration.

If the power was cycled during a 24 Hour program, the currently running program is restarted on power up using the appropriate step based on the current time.

### 24 Hour Ramping Program Type

The start time parameter in a 24 Hour program step is a time of day to begin using the programmed setpoints and light settings.

The end time is the start time parameter in the next chronological step.

The 24 Hour Ramping type starts with the programmed setpoints and dimming light settings at the start time of the step and ramps over the duration of the currently running step to the programmed setpoints and dimming light settings for the next step.

When a 24 Hour program is started, the first step is the one where the current time falls within the step duration.

When a 24 Hour Ramping program is started, the starting step setpoints and dimming lighting settings are calculated for the current time by using the start and end values and the duration of the step.

If the power was cycled during a 24 Hour program, the currently running program is restarted on power up using the appropriate step based on the current time.

### Elapsed Non-Ramping Program Type

The time parameter in an Elapsed program step is the duration to use the programmed setpoints and light settings.

The end time is determined by adding the duration to the time that the step started.

The Elapsed type uses the programmed setpoints and light settings for the duration of the currently running step.

When an Elapsed program is started, the first step is the first one in the program.

If the power was cycled during an Elapsed program, the currently running program step is restarted at the time the power went down.

### Elapsed Ramping Program Type

The time parameter in an Elapsed program step is the duration to use the programmed setpoints and light settings.

The end time is determined by adding the duration to the time that the step started.

The first step of the Elapsed Ramping type starts with the current Manual setpoints and dimming light settings at the start of the step and ramps over the duration of the currently running step to the programmed setpoints and dimming light settings of the currently running step.  Thereafter it starts with the previous step’s setpoints and light settings and ramps to the current step’s setpoints and light settings.

When an Elapsed program is started, the first step is the first one in the program.

If the power was cycled during an Elapsed program, the currently running program step is restarted at the time the power went down.

## Program Create/Edit Sequence

There can only be one sequence.

Create Sequence will not be available if no program has been created.

Edit Sequence will not be available if the sequence has not been created.

A Sequence consists of up to 50 steps containing a program name and the number of iterations for that program.

## Program Run Sequence

The sequence can be started from any step and also can be started at any iteration for that step.

When a sequence starts it starts the selected program as defined by the rules for starting a program.

When a sequence runs it uses the current step’s programmed iterations to determine how many times to cycle through the current program.

If a sequence was started at an iteration that is not 1, the program will cycle for that many less iterations.

When a sequence finishes its last step, the sequence begins again at the first iteration of the first step in the sequence, running through the entire sequence again.  This cycle continues until the Program Mode is changed.

### 24 Hour Non-Ramping Program Type

For a 24 Hour program type, an iteration is defined as the portion of the program that runs from 12.00 am to 11:59pm, so if a program is started in the middle of the day, the first iteration will complete at midnight.

For a 24 Hour Non-Ramping sequence, for the last iteration of the program which ends at midnight, the next step will be the program step that starts closest to midnight from the program in the next sequence step.

### 24 Hour Ramping Program Type

For a 24 Hour program type, an iteration is defined as the portion of the program that runs from 12.00 am to 11:59pm, so if a program is started in the middle of the day, the first iteration will complete at midnight.

For a 24 Hour Ramping sequence, for the last iteration of the program which ends at midnight, the next step will be the program step that starts closest to midnight from the program in the next sequence step.  This will be used to determine the end point for ramping the setpoints and lighting settings.

### Elapsed Non-Ramping Program Type

For an Elapsed program type an iteration is defined as the time from the start of the first step to the end of the last step.

For an Elapsed Non-Ramping sequence, for the last iteration of the program which ends at the last step, the next step will be the first step of the program in the next sequence step.

### 24 Hour Non-Ramping Program Type

For an Elapsed program type an iteration is defined as the time from the start of the first step to the end of the last step.

For an Elapsed Ramping sequence, for the last iteration of the program which ends at the last step, the next step will be the first of the program in the next sequence step.  This means that for the first step of the program in the next sequence step, the last step of the previous program will be used as the previous step for ramping.

## Program Edit Diurnal

Edit Diurnal will always be available even if Diurnal is already running.

Day and Night setpoints will be available for editing for any active Control Loops.

The Day Start Time and Day End Time parameters will be used to define Day and Night periods.

The Day Events Powered parameter will be used to define whether any Digital Ouputs set to Event On/Off will be turned on during a Day period.

## Program Run Diurnal

The Diurnal program mode defines Day as the time between the programmed Day Start Time and Day End Time.  The other period is defined as Night.

During the Day period, the Diurnal program mode uses the Diurnal Day setpoint for the control loops, uses 100% output power for all Digital Outputs programmed as Light On/Off, Light Dimming or DALI Dimming and uses the Day Events Powered parameter to decide whether the Digital Outputs programmed as Event On/Off are powered.

During the Night period, the Diurnal program mode uses the Diurnal Night setpoint for the control loops and uses 0% output power for all Digital Outputs programmed as Light On/Off, Light Dimming, DALI Dimming or Event On/Off.

## Program Edit DLI

Edit DLI will only be available if the Sensor Type for the selected DLI Input Source is set to a process type and the DLI Enable parameter is set to Yes.

Day and Night setpoints will be available for editing for any active Control Loops except the selected DLI channel will only show a Day setpoint (labled Day Lights) if the channel is configured as a control loop (really only possible with the RM channels).

The Day Start Time, Day End Time and Accrued DLI parameters will be used to define Day and Night periods.

The Day Events Powered parameter will be used to define whether any Digital Ouputs set to Event On/Off will be turned on during a Day period.

## Program Run DLI

The DLI program mode defines Day as the time between the programmed Day Start Time and Day End Time while the Accrued Daily Light Integral is less than the programmed Daily Light Integral.  The other period is defined as Night.

The Accrued Daily Light Integral is based on the level of the selected DLI Input Source in units of micromoles/meter2/second and is accrued every second.

The DLI program mode operates exactly the same as the Diurnal program mode with the exception that the Accrued Daily Light Integral value can cause the Day period to end before the Day End Time is reached.  Also, if the selected DLI channel is configured as a control loop the outputs will use the programmed Day Setpoint to control the loop during the day and will use 0°C as the setpoint during the Night.

The Accrued Daily Light Integral will only be cleared when the time period is night as defined by the time between Day End Time and Day Start Time.

# WeatherEze

The WeatherEze program is an external method for controlling an Intellus while the Intellus is running a Manual Program.

The WeatherEze program is expected to be controlling the setpoints and outputs when the WeatherEze Enable parameter is set to Yes and the Intellus is running a Manual Program.  (NOTE: If the WeatherEze parameter is set to Yes and there is no external program, the setpoints and outputs will be at their default values).

If the Intellus is not running a Manual Program, the WeatherEze externally supplied parameters are ignored.

When WeatherEze is enabled and a Manual Program is running, any active Control Loops will use the externally supplied setpoints.

The existing external program only controlled the first 16 outputs so when WeatherEze is enabled and a Manual Program is running, any of the first 16 Digital Outputs that are set to Light On/Off, Light Dimming or DALI Dimming will use the externally supplied power level.

# XML Tags

The XML tags are used in various features in the Intellus Gen II controller.  They are used in the communications over Ethernet, Bluetooth and USB.  In addition, the XML tag list is used for the auto configuration feature.

The complete XML tag list is attached:

## Format of Data Returned

In general the format of the data returned matches the format of the Gen I Intellus for the tags which are legacy.

There is one exception which was not discovered until the tablet application was far enough along to make changing it difficult.

For the temperature related parameters, all writes are required to be in integer form with implied three decimal places.  The reads are returned as floating point number in the precision of the related input or control loop.

## Encrypted XML Tags

There are two XML tags that require encryption when being read from the control.  These are the ControlOption1 and ControlOption2 tags which return the Admin User Password and the Monitor User Password that are used on the web pages.

The algorithm was provided by the customer, who also provided a tool (below) to verify that the algorithm is implemented correctly.

## HTTP and USB Command Strings

These command strings are used for the Ethernet, USB and Bluetooth communications protocols.  The XML Text referenced is from the XML Tag List document.  All Ethernet command strings should be prefaced by http://XXX.XXX.XXX.XXX/ where the Xs represent the unit’s IP address.

# Auto Configuration

The Auto Configuration feature allows storage to a USB memory stick of a list of parameters and values from an Intellus that has been configured.  The file name of the configuration file is created from the Chamber Serial Number parameter and has a .cfg extension.

The Auto Configuration feature allows retrieval from a USB memory stick of a list of parameters and values to an Intellus to configure it.  The file name of the configuration file is always “IntellusSeries3.cfg”.  If the USB memory stick does not have a file with this name on it in the root directory, the Configuration Load prompt will not be visible.

The list of parameters is defined by the XML Tag list.

# Datalogging

The Datalogging feature has two implementations depending on whether the Intellus is high end or low end.

A low end Intellus has a simple 24 hour volatile log that is used by the tablet for graphing.

A high end Intellus has a log that is saved on the SD card and that wraps after a year.

## Low End RAM Datalog

The low end Datalog information is stored in RAM every 5 minutes and wraps every 24 hours.

The parameters that are saved are Date, Time, Process Value for Analog Inputs 1-3, Setpoints for Control Loops 1-3, number of outputs programmed as lights, number of programmed lights that are on and the Limit 1 & 2 alarm setpoints high and low.

## High End SD Card Datalog

The high end Datalog information is stored on the SD card every 1 minute and wraps every year.

All parameters on the High End Datalog list are saved to the SD card.

The Datalog parameters can be retrieved from the SD card and saved to a USB memory stick.  The file name is created from the Chamber Serial Number and the current date in the Intellus and has a .csv extension.  A subset of the parameters on the list are saved to the USB file, the parameters saved depend on the configuration of the Intellus (i.e. Prog_Name would only be saved if a program or sequence was running).

### High End Datalog Parameter List

(These are the header strings for each column of the datalog file)

Year

Month

Day

Hour

Minute

Second

Mode

Prog_Name

Prog_Step

Seq_Step

Seq_Iter

PV_X	X=1-7

SP_X	X=1-7

EO_X	X=1-47

Defrost

DLI

DLI_Acc

Lim_1_Hi

Lim_1_Lo

Lim_2_Hi

Lim_2_Lo

Dev_Hi

Dev_Lo

Alm_X_Hi	X=1-7

Alm_X_Lo	X=1-7

Life_Alarm

Lim_1_Alarm

Lim_2_Alarm

Dev_Alarm

Alm_X_Alarm	X=1-7

Light_Life

ChXA_Ena	X=2-7

ChXB_Ena	X=2-7

LightX_Life	X=1-47

# Flashloading

The Intellus is designed to be able to reprogram the application code using a USB memory stick.  The file that is used is named PercivalAppArm.bin and it is created by the PostProcess.bat routine after the project is built.  This file is required to be in the root directory of the memory stick in order for the prompts to be displayed.

The Intellus flashloader has built in checking that allows restoring the previous version of the application if the new application does not run correctly.

## Successful Flashload Process

Insert the USB memory stick with the PercivalAppArm.bin file on it into the USB connector.

Press the Down Arrow key and the Diagnostics key for 3 seconds to move to the Firmware Download menu.

Verify the reported software revision and select Yes to start the flashloading process.

During this portion of the process the binary file is being stored in nonvolatile memory on the Intellus but the application is not being overwritten.

The display will show “Firmware Loading” on the top display and “Flashing – Erase”, "Flashing - Program", "Flashing - Verify" on the bottom display while the file is being transferred.

If the file transfer was not successful the bottom display will change to "Failed Flashload" and the process will be complete without trying to rewrite the application.

If the file transfer was successful the bottom display will change to "Successful Flashload" and after about 15 seconds the controller will reboot.

The process of rewriting the application is handled by the boot code after the reboot.

During the rewriting process the display will show “BOOT FLASH 5.0” on the top line and “WAIT” on the bottom line.  If the process is proceeding normally, there will be periods added to the display after the “WAIT” to show the progress.

When the rewrite is done, the controller will reboot again and transition to the new application.

## Corrupt Application After Flashload

The Intellus is designed to safely handle the case of a successful flashload of an application that does not run correctly.

After a successful flashload and application rewrite is completed, if the controller resets during the first 15 minutes of running, the boot code will assume that the new application is corrupt and will restore the previous version of code which is assumed to be safe.

During the rewriting process of the backup application the display will show “BOOT FLASH 5.0” on the top line and “WAIT” on the bottom line.  There will be periods added to the display after the “WAIT” to show the progress but in this case they do not necessarily show up in order and there is a X as one of the characters.

When the rewrite is done, the controller will reboot again and transition to the restored application.

# Secondary Micro

The Secondary Micro is intended to be a second limit in the system with its own sensor and its own control over the alarm buzzer and the power relay.

The information related to the Secondary Micro that is shown on the display uses the input name of Limit 2.

The Secondary Micro does not use nonvolatile storage for the parameters, the Intellus must send all the relevant parameter information at the initialization.

The Limit operation uses the Limit 2 Process Value and the programmed parameters Limit 2 Setpoint High and Limit 2 Setpoint Low.

## Secondary Micro Communications

The Secondary Micro and the Intellus application communicate using a serial UART.

### Intellus to Secondary Micro

The Intellus transmits a heartbeat or a parameter update message to the Secondary Micro every 100 mS.

The parameter update message overrides the normal heartbeat message if a Secondary Micro parameter has changed on the Intellus.

If a receive error is detected, an error message is sent to the Secondary micro within 10 mS.

### Secondary Micro to Intellus

The Secondary Micro transmits a status message to the Intellus application every 100 mS.

If a receive error is detected, an error message is sent to the Intellus application within 10 mS.

## Secondary Micro Analog Input

This input is always used for temperature sensing using a 100 ohm RTD sensor and it is always enabled.

The units for this input follow the display units setting, either F or C.

## Secondary Micro Buzzer Control

The Secondary Micro has its own separate control of the buzzer output.

The buzzer output is turned on if a sensor error is detected or the process value is above the Limit 2 High Setpoint or below the Limit 2 Low Setpoint.

Pressing any key will cause the buzzer output to be turned off.

## Secondary Micro Power Relay Control

The Secondary Micro has its own separate control of the power relay.

The power relay will be turned on if the communications between the Secondary Micro and the Intellus has been active for the previous 10 seconds.

The power relay will be turned off if a sensor error is detected or the process value is above the Limit 2 High Setpoint or below the Limit 2 Low Setpoint or there has not been a heartbeat message from the Intellus for 10 seconds.

## Secondary Micro Reprogramming

The Intellus application code is designed to reprogram the Secondary Micro automatically if the software revision of the Secondary Micro code does not match the revision of the image which is part of the application code or if there is no response on power up.

# Security

Security is disabled if the three security level passwords are set to their default value.

If one or more of the passwords are changed, security is enabled and menu access will be restricted unless the user is logged in with a correct password.

The security level passwords are available in the User/Global menu.

## Logging In and Out

In the display loop pressing the Down Arrow key and the Alarms key for 3 seconds will move to the Login/Logout menu if security is enabled.

Entering a correct password in this menu will log the user in with the Security Level of the password they have entered.

A user will automatically be logged out if they are logged in and 5 minutes have elapsed with no key presses.

If a user is logged in, they will be able to choose to Logout by reentering the Login/Logout menu as stated above.

## Security Level 4

Security Level 4 is the most restrictive and it is active when security is enabled and there is no one logged in.

It allows access only to the prompts in the display loop and to the password protected menus if the correct password is entered.

## Security Level 3

The default password for Security Level 3 is “0003”.

Security Level 3 allows access to everything accessible at Security Level 4 plus:

- User/Alarm menu, Alarm High Setpoint or Alarm Low Setpoint only.

- User/Diagnostics menu.

- Service Diagnostics menu.

## Security Level 2

The default password for Security Level 2 is “0002”.

Security Level 2 allows access to everything accessible at Security Levels 3 and 4 plus all of the other User menus.

## Security Level 1

The default password for Security Level 1 is “0001”.

The hard coded password of “SLVL” will also always select Security Level 1.

Security Level 1 allows access to everything as if security was disabled.

## Menu Passwords

There are two menus that require passwords to access, whether or not security is enabled.

In the display loop pressing the Up Arrow key and the Enter key for 3 seconds will move to the password protected User\Global menu.

There are two passwords that will allow access to the User/Global menu. “SLVL” is a hard coded backdoor password.  The other password is available for editing in the User/Global menu.  It is defaulted to “1111”.

In the display loop pressing the Up Arrow key and the Down Arrow key for 3 seconds will move to the password protected Setup menus.

There are two passwords that will allow access to the Setup menus.  “PICS” is a hard coded password that will allow access to all the Setup menus.  “APID” is a hard coded password that will allow access to only the Setup/PID menu.

## Password Override

During Power Up, pressing the Enter key and the Alarms key will enable the password override until the power is cycled.

The password override is a method for disabling the requirements for entering a password to access any of the menus.  This is intended to be used for the customer to more easily do the configuration of the control.

If the password override is enabled, the key timeout which transitions from a menu to the display loop is disabled.

# DALI

The DALI communications feature is intended to be used with special ballasts on fluorescent lamps to do light dimming control.

The normal light dimming feature outputs a PWM signal on the selected output.  The DALI Dimming feature transmits the expected light dimming power level using the DALI communications to the selected ballast address and sets the physical output to off.

There are a total of 64 possible addresses on the DALI bus.

The DALI Configuration is done in the Setup\DALI Configuration menu.

The first prompt chooses the method of transmission, Off, Broadcast or Direct.

The next prompt(s) select the source for the transmitted power level.

The last prompts allow programming an address into a DALI ballast.

## Transmission Method Off

If the Transmission Method is set to Off, no transmission is sent to any of the addresses.

## Transmission Method Broadcast

If the Transmission Method is set to Broadcast, only one source is allowed to be selected in the DALI Configuration.  This source’s power level is used in a broadcast command to any DALI ballasts on the bus, no specific addresses are used.

## Transmission Method Direct

If the Transmission Method is set to Direct, each of the 64 addresses can be individually set to have a specific source.

Also, any of the addresses can select the same source, which allows “grouping” through programming the sources rather than by using the grouping mode on the DALI ballasts.

## Power Sources

The transmitted power level can be sourced from the Control Loop 3 (Auxiliary) Channel A output or from any Digital Output that is configured as a DALI Dimming type.  The Power Source can also be set to Off which always transmits 0% power.

NOTE: There is no checking in the DALI configuration as to whether the selected digital output is actually set to the DALI Dimming type.  If a Power Source is set to a Digital Output which is NOT configured as a DALI Dimming type, 0% power will always be transmitted to the DALI address.

## Programming a DALI ballast address

Individually programming a DALI ballast requires that only one ballast be attached to the DALI bus during the programming.

The first prompt after the Power Source is selected allows setting the address that is to be programmed.

The second prompt requests the address programming to the DALI ballast.

# Battery Backed RAM Parameters

There are certain parameters that need to be written so frequently that they need to be saved in the battery-backed RAM in the Real Time Clock module.

- Power Loss Time – used for time calculations using the length of time of the power loss.

- Deviation Alarm Latching Active – used to prevent a user from cycling the power to clear a latched deviation alarm.

- Active Program Mode – used to allow the control to run in the same mode after a power cycle.

- Active Program or Sequence Information, File, Step, Iteration, Step Time – used to allow restarting a program after a power cycle.

# Modbus Building Automation System

This feature uses a standard UART interface to an RS-485 bus to provide external systems with control information through a Modbus protocol.

The two parameters that are available to configure the Modbus port are the address and the baud rate and they are found in the Setup/Global menu.  The address can be set from 1-247, the baud rate can be set to 9600, 19200 or 38400 baud.

The parameters that are available to be read are those that are on the High End Datalog list.

The attached file shows the register numbers and their corresponding parameter.

# RM Communications

The Intellus Gen II control communicates to the external RMs through a High Speed Watbus interface.

There can only be two external RMs attached, one RMC and one RME.

The RMC is always assumed to be set to Zone 2 and the RME is assumed to be Zone 3.  Any other settings will not be recognized by the control.

If there is no RM recognized for the particular zone the features are set to their disabled modes.  For the RMC, AIN4-7 Sensor Type are set to Off, Control Loops 4-7 Control Modes are set to Off, Channels A and B outputs are disabled and Soft Alarms 4-7 are disabled.  For the RME, Digital Output Type is set to Off, Digital Input Action is set to Off and Analog Output Source is set to Off.

# USB Communications

The tablet is intended to be plugged into the USB port closest to the PCB since that port is the only one configured for charging.

The USB stick is intended to be plugged into the USB port farthest from the PCB.

The dual stacked USB ports are used for the following features on the Gen II Intellus:

- Flashloading new code from the USB stick

- Transferring Auto Configuration files from the Intellus to the USB stick or from the USB stick to the Intellus

- For the High End units only, saving Datalog files from the Intellus to the USB stick

- Tablet connection with communications using the XML tag protocol for USB

# Ethernet

The Ethernet hardware is only installed on High End units and the functionality is disabled on Low End units.

The Ethernet connection is used to receive external commands using the http protocol and the XML tags, serve web pages, push data to the Cloud and send e-mails.

The parameters related to the mode and address of the Ethernet connection are found in the User/Global menu.  The method for determining the IP address can be set to DHCP/Fixed or Fixed.  If it is set to DHCP/Fixed and the DHCP cannot be established, the method will automatically switch to Fixed.

The connection has auto detect for the speed of 10Mbps or 100Mbps and also for full or half duplex.

The Net Bios name service is installed to allow accessing the Intellus with the device name instead of the IP address.

The web page connection is unsecure except for the Setup page.  The http protocol using XML uses the unsecure connection.  The e-mail connection is always secure.  The Cloud connection can select between secure and unsecure.

# Cloud Connection

The Cloud connection feature is only available on High End units.  The data samples are taken from the SD card, so an SD card must be installed for this feature to work properly.

The parameters related to the mode and address of the Cloud connection are found in the User/Global menu.

Cloud Server IP Address	Read/Write	IP Address String (e.g. 192.168.1.1)

Cloud Connection Type 	Read/Write	Non-SSL / SSL

Cloud Default Port	Read/Write	0-65535
Cloud Secure Port	Read/Write	0-65535
Customer Identifier	Read/Write	8 characters (0-9, A-Z, a-z)

Unique Controller ID	Read		Zero padded hex MAC address (fixed width hex, no ‘:’s or ‘-‘s)

Connection Enabled	Read/Write	TRUE / FALSE

If any of the above parameters are changed, the controller forces an auto disconnect/reconnect sequence.

The session identifier is provided by the Saturn app and receipt identifier is an incrementing counter.  Both are eight hexadecimal characters.

If the connection type is set to SSL the connection is secure, the certificate is used and the secure port number is used.  Otherwise the default port is used.

This feature only pushes data from the Intellus to the Cloud.  The parameters that are transferred are the same set that are saved for the High End Datalogging and the data format is the same as if it were being stored in the USB datalog file.

The header transmits the strings used by the USB datalog file except that it only transmits the string for the digital output power (used for Dimming), skipping the string for the digital output state (used for On/Off) because the text was the same for both and it is duplicate information for the outputs.  Other configuration information is appended to the header on separate lines for each type (see attached document for details).

The attached document is produced by the tablet application vendor to describe the requirements for the Intellus to interface to the Cloud.

# Bluetooth

The Bluetooth hardware is only installed on High End units and the functionality is disabled on Low End units.  The Gen II Intellus implements a custom multi-point bluetooth comms link with a Google Nexus tablet via Serial Port Profile (SPP).  Once a client device is paired with the Intellus, the requests and responses are in the XML format.

# Web Pages

The web pages are only available on the High End units.  They are stored in the Http directory on the SD card.

Screen shots of the web pages are shown in the attached document.

## Web Page Security

There are two levels of password security in the web pages, Administrator and General User.

### Administrator Security

The default Administrator User Name is “ADMIN”, the default password is “PSIADMIN”.  These parameters can only be changed by a user who has entered the Administrator user name and password correctly.

The backdoor password for the Administrator security is User Name “always” and Password “works”.

### General User Security

The default General User User Name is “USER” and the default General User Password is blank.  The General User name and password cannot be used until the Administrator sets the user name and password.

The General User will have access to all the web pages except the Setup page.

# Email

The Email feature requires the Ethernet hardware which is only installed on High End units, so the functionality is disabled on Low End units.

The Email requires a secure connection and the Intellus Gen II has certificates for gmail.com, yahoo.com and mail.com.

There are three different kinds of e-mails that can be sent, Test, Cyclic and Change of State.  None of these are exclusive, Test is a one-time request, but both Change of State and Cyclic could be set active at the same time.

## Test

The Test email request is triggered by checking a box on the web page where the e-mail feature is configured.

The text of the e-mail will have the format as shown:

E-mail Alert Test

Device name: PSIxxxxx

This is a test of the e-mail alert system.

## Cyclic

The Cyclic email request is based on a timed interval which can be set between Off, 1 hour, 6 hours, 12 hours and 24 hours.  The content of the e-mail depends on what is enabled on the Intellus, i.e. if an alarm is disabled the alarm setpoints will not be part of the e-mail.

The text of the e-mail will have the format as shown:

Current Conditions, Controller: PSIxxxxxx

Controller Time: 10:13 AM

Operation: Manual

Actual TEMP: 38.2 F

TEMP Set Point: 75.2 F

High Limit 1 Alarm Set Point: 131.0 F

Low Limit 1 Alarm Set Point: 23.0 F

High Limit 2 Alarm Set Point: 140.0 F

Low Limit 2 Alarm Set Point: 17.6 F

High TEMP Alarm Set Point: 300.2 F

Low TEMP Alarm Set Point: -50.8 F

Actual RH  : 0 %RH

RH   Set Point: 50 %RH

Actual AUX : -0 xZX

AUX  Set Point: 0 xZX

High AUX  Alarm Set Point: 300 xZX

Low AUX  Alarm Set Point: -51 xZX

Actual RTD4: 1039 RM1

RTD4 Set Point: 0 RM1

Actual LS5 : 9 RM2

Actual RTD6: 0.0 RM3

Actual RTD7: 9.1 RM4

Light 1:On

Light 2:Off

Light 36:On

Light 37:On

Light 38:Off

Light 39:On

Light 40:On

Light 41:On

Light 42:On

Light 43:On

Light 44:Off

Light 45:Off

Light 46:Off

Light 47:On

## Change of State

The Change of State email request is only caused by an alarm becoming active.  Any of the analog inputs (1-7) which are enabled and have either a soft alarm or a sensor error will trigger a request.  In addition, the two limit alarms, the analog input 1 deviation alarm and the light lifetime alarm will trigger a request.

The text of the alarm portion of the e-mail will have the format as shown, with the text from the cyclic status being added after the alarms:

Current Alarms / Errors, Controller: PSIxxxxxx

IMPORTANT! - Please investigate any errors / alarms at the chamber.

Alarm: E01 Sensor Error

Alarm: E02 Sensor Error

Alarm: E03 Sensor Error

Alarm: E04 Sensor Error

Alarm: E05 Sensor Error

Alarm: E06 Sensor Error

Alarm: E07 Sensor Error

Alarm: E20 Sensor Error

Alarm: Light Lifetime

# Real Time Clock

The realtime clock is battery backed up to keep the time through power cycles.  When a power up occurs, the current time is retrieved and saved for diagnostic purposes.  The current time is used in many of the features.